## Slide 1
9th Renewable Energy Postgraduate Symposium (REPS)
Stellenbosch University
13 – 14th September, 2018

## Slide 2
Irradiance Feedback Control System for a 5-kW Xenon Arc Lamp Solar Simulator
Julian C. Nwodo
Supervisor: Prof. Edson L. Meyer
Co-Supervisor: Dr. Richmore Kaseke
Co-Supervisor: Dr. Raymond T. Taziwa
13th September, 2018

## Slide 3: Outline
Introduction
Aim
Objectives
Background
  Materials and Methods
Hardware
  System Architecture
Methodology
  Results
  Lamp Startup
  Irradiance Non-Uniformity
  Irradiance and Temperature Test
  Conclusions
  References
  Acknowledgements

## Slide 4: Introduction
The combination of a number of equipment or components connected in a particular order to perform a specific task is often referred to as a system.
When the system’s input quantity is controlled by varying the input quantity then the system is called as a control system.

## Slide 5: Introduction
The control system should regulate or direct or command.
Hence, a control system is an arrangement of distinct physical components connected in such a manner so as to:
Regulate or direct or to direct or to command itself or some other system.

## Slide 6: Introduction
Types of Control Systems
Open loop control systems
Closed loop control system
Open loop control systems
PROCESS
OUTPUT
INPUT

## Slide 7: Introduction
Types of Control Systems
Closed loop control systems
Process
INPUT
OUTPUT
FEEDBACK
Controller

## Slide 8: Aim and Objectives
Aim:
To design an irradiance feedback control system for a 5-kw xenon arc lamp solar simulator.

## Slide 9: Aim and Objectives
Objectives:
Design of an irradiance feedback control system.
Minimize human-machine interaction while carrying out indoor tests.
To achieve solar simulator set point quicker thereby reducing temperature build up on the target area.

## Slide 10: Materials and Methods
System Architecture
Figure 1: Schematic diagram of the irradiance feedback control system

## Slide 11: Materials and Methods
Hardware Description
Power Supply Units
MCB
MOV
Terminal Blocks
DIN Rail
Relays
TVS
PLC
DAIO
8TCISO
Ethernet Cable
TCP/IP Port
Figure 2: Schematic diagram of the irradiance feedback control system

## Slide 12: Materials and Methods
Methodology

## Slide 13: Methodology

## Slide 14: Methodology

## Slide 15: Methodology
Irradiance, Temperature & Non-Uniformity Test
Target area 2m x 2m
Figure 4: Target area

## Slide 16: Device Under Test
Reference cell
Pyranometer
Spectroradiometer
Figure 5: Target area showing device under test
Adjustable Unistrut Frames

## Slide 17: METHODOLOGY
Characterisation
Chart: Wavelength (mm)
5 kW DC Power Supply
Igniter
5 kW Xe Arc Lamp
Reflector with Xe-Arc Lamp
Auto/Manual Lamp Control
Target Area with Reference cell, Pyranometer and Spectroradiometer

## Slide 18: METHODOLOGY
Applying the Xenon-lamp solar simulator on various PV technologies
Chart: Voltage (V)
5 kW DC Power Supply
Igniter
5 kW Xe Arc Lamp
Reflector with Xe-Arc Lamp
IFCS
Target Area with Reference cell and C-Si Module

## Slide 19: Results
  Lamp Startup
Figure 4: Xenon arc lamp power and irradiance distribution
Figure 6: Xenon arc lamp power and irradiance distribution
Irradiance 67 Wm-2
Power 600 W
Lamp stability 5 sec
± 1% drift from set point

## Slide 20: Results
Figure 7: Solar simulator operated at varying set points of its rated current
Varying Lamp Input Power and Temperature Test
1000 Wm-2 at 90% max current, 20 s, 35 ℃
± 2.3% irradiance drift from set point

## Slide 21: Results
Table 1: Non-Uniformity at varying irradiance set points
Non-Uniformity
Power level | Current Drawn (A) | Cell surface Temperature (℃) | Non-Uniformity (%)
20 % of rated max | 29 | 22.6 | 1.02
50 % of rated max | 72.5 | 24.4 | 1.53
90 % of rated max | 130.5 | 25.4 | 3.26

## Slide 22: Conclusions
1000 Wm-2  90%  <10 s to stabilize…
1000 Wm-2  3.26%  Class A
Achieves Set points faster,…
IRRADIANCE
Non-Uniformity
PERFOMANCE

## Slide 23: References
www.studymafia.org [Accessed on 05/09/2018]
https://www.indiamart.com [Accessed on 05/09/2018]
https://www.banggood.com [Accessed on 05/09/2018]
http://www.tvms.co.za [Accessed on 05/09/2018]
https://solanoheating.com [Accessed on 05/09/2018]
Rivola, D., Dittmann, S., Pravettoni, M., Friesen, G. and Chianese, D., 2014, June. High-speed multi-channel system for solar simulator irradiance non-uniformity measurement. In Photovoltaic Specialist Conference (PVSC), 2014 IEEE 40th (pp. 2611-2615). IEEE. DOI: 10.1109/PVSC.2014.6925465
Chawla, M.K. and Tech, P.E., 2018. A step by step guide to selecting the “right” Solar Simulator for your solar cell testing application. Photo Emission Tech., Inc. simulator for your solar cell testing application.” <http://www.photoemission.com/techpapers/A%20step%20by%20step%20guide%20to%20selecting%20a%20Solar%20Simulator%20Ver.%203.pdf> (Retrieved 29.03.18).
Pravettoni, M., Galleano, R., Aitasalo, T., Kenny, R.P., Dunlop, E.D. and Barnham, K.W., 2010, June. From an existing large area pulsed solar simulator to a high intensity pulsed solar simulator: characterization, standard classification and first results at ESTI. In Photovoltaic Specialists Conference (PVSC), 2010 35th IEEE (pp. 002724-002728). IEEE. DOI: 10.1109/PVSC.2010.5616862
Bazzi, A.M., Klein, Z., Sweeney, M., Kroeger, K., Shenoy, P. and Krein, P.T., 2011, March. Solid-state light simulator with current-mode control. In Applied Power Electronics Conference and Exposition (APEC), 2011 Twenty-Sixth Annual IEEE (pp. 2047-2053). IEEE. DOI: 10.1109/APEC.2011.5744878
Makosa, C., Meyer, E. L., Nwodo, J. C., Kaseke, R. and Taziwa, R. T. 2017. Characterization and Classification of a 5-kW Xenon Lamp Solar Simulator with an Ellipsoidal Reflector. SAIP Conference 3rd – 7th July 2017.

## Slide 24: Acknowledgements
Prof. E. L. Meyer
Dr. R. Kaseke
Dr. R. T. Taziwa
M. Adonis

## Slide 25
THANK YOU FOR LISTENING